## Slide 1
National Voluntary Laboratory Accreditation Program
C63 SC6
October 2, 2024

## Slide 2: Overview
Number of organizations accredited to C63 standards
Top nonconforming issues to C63 standards
Top nonconforming issues to accreditation standards and AB requirements
Summary of AB Accreditation Schemes using C63 standards (i.e., 17025, 17065, etc.) and associated programs (FDA, FCC, ISED, etc.)

## Slide 3: Number of organizations accredited to C63 standards
Standard | May 2024 | October 2024 | May 2025
ANSI C63.4 (2014) | 61 labs | 60 labs | 60 labs
ANSI C63.10 (2013) | 39 labs | 51 labs | 39 labs
ANSI C63.10 (2020) | 18 labs | 19 labs | 22 labs
ANSI C63.17 (2013) | 14 labs | 14 labs | 14 labs
ANSI C63.19 (2011) | 1 lab | 1 lab | 1 lab
ANSI C63.26 (2015) | 28 labs | 31 labs | 24 labs

## Slide 4: Top Non-conformities to C63 Standards
1. ANSI C63.4-2014, clause 4.7.3 Antenna calibration
All antennas shall be calibrated in accordance with ANSI C63.5, including antennas used for performing NSA measurements for standard test site validation and alternative test site validation (see 5.4.2) and antennas used for performing SVSWR measurements for test site validation above 1 GHz.
2. ANSI C63.4:2014, clause 4.2.3C
No provisions in place to validate that linearity test was performed per the requirement of the standard.
3. ANSI C63.4-2014, clause 10.2.4 "List of test equipment
A complete list of all test equipment used shall be included with the test report. The manufacturer’s model and serial numbers, date of last calibration, and calibration interval shall be included. Measurement cable loss, measuring instrument bandwidth and detector function, video bandwidth, if appropriate, and antenna factors shall also be included where applicable."

## Slide 5: Top Non-Conformities to 17025
1. Missing procedures and/or records (6.2.5)
  Determining the competence requirements
  Selection of personnel
  Training of personnel
  Supervision personnel
  Authorization of personnel
  Monitoring competence of personnel
2. Equipment Records (6.4.13)
3. Measurement Uncertainty (7.6.1)
4. Reporting of Results (7.8.1.2)
5. Common Requirements for Reporting (7.8.2)

## Slide 6: Summary of Accreditation Schemes
 | May 2024 | October 2024 | May 2025
ISO 17025 – ECT Program | 88 | 87 | 84
ASCA | 9 | 8 | 6
ISED | 35 | 36 | 35
FCC | 46 | 47 | 47

## Slide 7: Contact Information
NVLAP
National Institute of Standards and Technology
100 Bureau Drive Stop 2140
Gaithersburg, MD 20899-2140
Amanda McDonald, ECT Program Management
Voice: +1.301.975.5627
Amanda.mcdonald@nist.gov
Janneth Marcelo, ECT Program Management
Voice: +1.301.975.8520
Janneth.Marcelo@nist.gov